Acylcarnitine LC/MS/MS Assay

Sanford-Burnham Medical Research Institute at Lake Nona, Orlando, FL

Chris Petucci, Ph.D. and Jeff Culver, M.S.

- Preparation of Calibration and Internal Standard Solutions

- Preparation of Individual Acylcarnitine Stock Solutions:  Individual acylcarnitine stock solutions from powders are prepared according to table below.

| Acylcarnitine (AC) | Carrier | Cat. No. | Individual Stock Conc. (M) | Diluent |
| --- | --- | --- | --- | --- |
| Acetylcarnitine, C2 | Sigma-Aldrich | A6706 | 0.01 | DI water |
| Propionylcarnitine, C3 | R&D Systems | 0611/50 | 0.01 | DI water |
| Butyrylcarnitine, C4 | Toronto Research Chemicals | B802700 | 0.01 | DI water |
| Isovalerylcarnitine, C5-Iso | Larodan | 17-0550-9 | 0.01 | DI water |
| Hexanoylcarnitine, C6 | R&D Systems | 0526/50 | 0.01 | DI water |
| Octanoylcarnitine, C8 | R&D Systems | 0605/50 | 0.01 | 50% MeOH |
| Decanoylcarnitine, C10 | R&D Systems | 0477/50 | 0.01 | 50% MeOH |
| Lauroylcarnitine, C12 | R&D Systems | 0548/50 | 0.01 | 50% MeOH |
| Myristoylcarnitine, C14 | R&D Systems | 0567/50 | 0.01 | 50% MeOH |
| Palmitoylcarnitine, C16 | Sigma-Aldrich | P4509 | 0.01 | 50% MeOH |
| Stearoylcarnitine, C18 | Larodan | 17-1800-9 | 0.01 | 50% MeOH |

- Preparation of Acylcarnitine Working Solutions:  Individual stock solutions are combined to prepare the highest calibrator working solution (C11 in the table below).

| Acylcarnitine | Initial Individual Stock Conc. (M) | Final Combined C11 Stock Conc. (mM) |
| --- | --- | --- |
| C2 | 0.01 | 0.5 |
| C3 | 0.01 | 0.05 |
| C4 | 0.01 | 0.05 |
| C5-Iso | 0.01 | 0.05 |
| C6 | 0.01 | 0.05 |
| C8 | 0.01 | 0.05 |
| C10 | 0.01 | 0.05 |
| C12 | 0.01 | 0.05 |
| C14 | 0.01 | 0.05 |
| C16 | 0.01 | 0.05 |
| C18 | 0.01 | 0.05 |

Serial dilution of this working solution in 50/50 Acetonitrile/Water with 0.3% Formic Acid yields all eleven calibrator working solutions, C1-C11 in the table below:

| Calibrator Stock Soln. | Conc. of C2 (μM) | Conc. of All Other ACs (μM) |
| --- | --- | --- |
| C11 | 500 | 50 |
| C10 | 250 | 25 |
| C9 | 100 | 10 |
| C8 | 50 | 5 |
| C7 | 25 | 2.5 |
| C6 | 10 | 1 |
| C5 | 5 | 0.5 |
| C4 | 2.5 | 0.25 |
| C3 | 1 | 0.1 |
| C2 | 0.5 | 0.05 |
| C1 | 0.25 | 0.025 |

A 10 µL aliquot of the calibrator working solutions above produces individual calibration curve concentrations used for LC/MS/MS quantitation of acylcarnitines according to the table below:

| Calibrator Stock Soln. | Conc. of C2 (μM) | Conc. of All Other ACs (μM) |
| --- | --- | --- |
| C11 | 50 | 5 |
| C10 | 25 | 2.5 |
| C9 | 10 | 1 |
| C8 | 5 | 0.5 |
| C7 | 2.5 | 0.25 |
| C6 | 1 | 0.1 |
| C5 | 0.5 | 0.05 |
| C4 | 0.25 | 0.025 |
| C3 | 0.1 | 0.01 |
| C2 | 0.05 | 0.005 |
| C1 | 0.025 | 0.0025 |

- Preparation of Individual Acylcarnitine Internal Standard Stock Solutions:  Individual acylcarnitine internal standard stock solutions are prepared from powders according to table below.

| Acylcarnitine-Int. Std. | Carrier | Cat. No. | Stock Conc’n (M) | Diluent |
| --- | --- | --- | --- | --- |
| Acetylcarnitine, C2-D3 | Cambridge Isotopes | DLM-1871 | 0.01 | DI water |
| Propionylcarnitine, C3-D3 | Cambridge Isotopes | DLM-3973 | 0.01 | DI water |
| Butyrylcarnitine, C4-D3 | Cambridge Isotopes | DLM-3861 | 0.01 | DI water |
| Isovalerylcarnitine, C5-Iso-D9 | Cambridge Isotopes | DLM-3974 | 0.01 | DI water |
| Hexanoylcarnitine, C6-D3 | CDN Isotopes | D-6647 | 0.01 | DI water |
| Octanoylcarnitine, C8-D3 | Cambridge Isotopes | DLM-755 | 0.01 | 50% MeOH |
| Decanoylcarnitine, C10-D3 | CDN Isotopes | D-6640 | 0.01 | 50% MeOH |
| Dodecanoylcarnitine, C12-D3 | CDN Isotopes | D-6643 | 0.01 | 50% MeOH |
| Tetradecanoylcarnitine, C14-D3 | CDN Isotopes | D-6661 | 0.01 | 50% MeOH |
| Palmitoylcarnitine, C16-D3 | Cambridge Isotopes | D-1263 | 0.01 | 50% MeOH |
| Octadecanoylcarnitine, C18-D3 | CDN Isotopes | D-6650 | 0.01 | 50% MeOH |

- Preparation of Acylcarnitine Internal Standard Working Solution I: Individual internal standard stock solutions are combined in 50/50 Acetonitrile/Water with 0.3% Formic Acid to prepare the internal standard stock working solution.

| Acylcarnitine-Int. Std. | Initial Individual Stock Conc. (M) | Final Combined Stock Conc. (μM) |
| --- | --- | --- |
| C2-D3 | 0.01 | 25 |
| C3-D3 | 0.01 | 2.5 |
| C4-D3 | 0.01 | 2.5 |
| C5-Iso-D9 | 0.01 | 2.5 |
| C6-D3 | 0.01 | 2.5 |
| C8-D3 | 0.01 | 2.5 |
| C10-D3 | 0.01 | 0.25 |
| C12-D3 | 0.01 | 2.5 |
| C14-D3 | 0.01 | 2.5 |
| C16-D3 | 0.01 | 0.25 |
| C18-D3 | 0.01 | 2.5 |

The internal standard working solution above is diluted 1:10 in 50/50 ACN/0.3% formic acid to generate the internal standard working solution II.

A 10 µL aliquot of internal standard working solution II is spiked into calibrator and study samples to produce the following concentrations for LC/MS/MS quantitation of acylcarnitines:

| Sample | Conc. of “A” ISs (uM) | Conc. of “B” ISs (uM) | Conc. of “C” ISs (uM) |
| --- | --- | --- | --- |
| C1-C9 | 2.5 | 0.25 | 0.025 |
| Tissue/Fluid | 2.5 | 0.25 | 0.025 |

“A” Internal Standard: C2
                               “B” Internal Standards: C3, C4, C5-Iso, C6, C8, C12, C14, C18 
                               “C” Internal Standards: C10, C16

- Preparation of Calibration Curves and Biological Samples

- Preparation of Calibration Curves

- To eleven Eppendorf tubes, add either 90 μL of either H2O (for fluid analysis) or 50/50 ACN/0.3% formic acid (for tissue analysis).

- Spike in 10 μL of the appropriate calibrator working solution to each tube.

- Spike in 10 μL of internal standard working solution II to each tube and vortex thoroughly.

- For all remaining steps, beginning with the addition of MeOH, follow the steps used in “Preparation of Biological Samples” below.

- Preparation of Biological Samples

- Aliquot 100 μL of either biological fluid or tissue homogenate (at a concentration of 100 mg tissue per mL of 50/50 ACN/0.3% formic acid) into an Eppendorf tube.

- Add 10 μL of internal standard working solution II to each tube.

- Add 800 μL of ice-cold MeOH to the tube and vortex thoroughly.

- Centrifuge the tube at 18,000 x g for five minutes at 4°C; precipitated protein will form a tight pellet at the bottom of the tube.

- After centrifugation, transfer 100 μL of methanolic supernatant to a 96-well plate.

- Dry the contents of the plate under nitrogen at 45°C until all wells are completely dry

- Meanwhile, prepare the derivatization reagents

  - Prepare 0.2 M o-benzylhydroxylamine (OBA) in 50% MeOH/50% 10 mM ammonium acetate, pH 5.5 (store at room temperature; this solution is stable for three days).

  - Prepare 2 M EDC in H2O (this solution must be prepared daily).

- Once the wells are dry, reconstitute the samples in 100 μL of OBA and 10 μL of EDC. Seal the plate, vortex on a plate mixer for 30 seconds, and incubate for 10 minutes at room temperature.

- After incubation, transfer the plate to the autosampler of the LC/MS/MS.

- LC/MS/MS Assay Conditions

- Agilent 1290 Infinity HPLC

Autosampler

  - Temperature:  4°C

  - Injection Volume:  2 μL

  - Needle Wash Solution:  80/20 Methanol/Water

Column

  - Waters 2 x 100 mm, 1.7 µm BEH C18 column

  - Temperature:  55°C

  - Maximum Pressure:  1000 bar

Binary Pump

  - Flow Rate:  0.4 mL/min

  - Solvent A:  0.1% Formic Acid in H2O

  - Solvent B:  0.1% Formic Acid in ACN

  - Gradient Conditions:

| Segment | Time (min.) | % B | Flow Rate (mL/min) |
| --- | --- | --- | --- |
| 0 (Start) | 0.00 | 1 | 0.4 |
| 1 | 3.00 | 20 | 0.4 |
| 2 | 5.00 | 30 | 0.4 |
| 3 | 12.50 | 95 | 0.4 |
| 4 | 12.60 | 95 | 0.8 |
| 5 | 14.10 | 95 | 0.8 |
| 6 | 14.20 | 1 | 0.8 |
| 7 | 16.00 | 1 | 0.8 |
| Re-equil. | 16.20 | 1 | 0.4 |

- Agilent 6490 Triple Quadrupole Mass Spectrometer

  - Gas Temperature: 325˚ C

  - Gas Flow: 11 L/min.

  - Nebulizer: 50 psi

  - Sheath Gas Temperature: 325˚ C

  - Sheath Gas Flow: 10 L/min.

  - Capillary Voltage: 3500 V

  - Nozzle Voltage: 500 V

  - Electrospray ionization: Positive

- MRM Transitions

| Compound Name | Precursor Ion | Product Ion | MS2 Res | Dwell | Fragmentor | Collision Energy | Cell Accelerator Voltage |
| --- | --- | --- | --- | --- | --- | --- | --- |
| C22 | 589.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C22:1 | 587.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C22:2 | 585.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C22:3 | 583.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C22:4 | 581.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C22:5 | 579.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C20-OH/C22:6 | 577.3 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C20:1-OH | 575.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C20:2-OH | 573.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C20 | 561.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C20:1 | 559.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C20:2 | 557.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C20:3 | 555.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C20:4 | 553.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C18-OH | 549.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C18:1-OH | 547.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C18:2-OH | 545.2 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C18-d3 | 536.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C18 | 533.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C18:1 | 531.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C18:2 | 529.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C16-OH | 521.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C16:1-OH | 519.3 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C16:2-OH | 517.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C16-d3 | 508.2 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C16 | 505.2 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C16:1 | 503.2 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C16:2 | 501.2 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C14-OH | 493.2 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C14:1-OH | 491.2 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C14:2-OH | 489.2 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C5-DC | 486.2 | 190.1 | Unit | 3 | 380 | 30 | 5 |
| C14-d3 | 480.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C14 | 477.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C14:1 | 475.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C14:2 | 473.3 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| Ci4-DC | 472.2 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C4-DC | 472.2 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C12-OH | 465.1 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C3-DC | 458.1 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C12-d3 | 452.2 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C12 | 449.2 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C12:1 | 447.2 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C10-OH | 437.1 | 190.1 | Unit | 3 | 380 | 25 | 5 |
| C10-d3 | 424.2 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C10 | 421.2 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C8-OH | 409.1 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C8:1-OH | 407.1 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C8-d3 | 396.2 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C8 | 393.2 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C6-OH | 381.1 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C6-d3 | 368.2 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C5-OH | 367.2 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C6 | 365.2 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C5-d9 | 360.2 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C4-OH | 353.2 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C5 2-methylbutyryl | 351.2 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C5 isovaleryl | 351.2 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C5 valeryl | 351.2 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C5:1 | 349.2 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C4-d3 | 340.2 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C4 butyryl | 337.2 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C4 isobutyryl | 337.2 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C4:1 | 335.2 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C3-d3 | 326.2 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C3 | 323.2 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C2-d3 | 312.2 | 190.1 | Unit | 3 | 380 | 20 | 5 |
| C2 | 309.2 | 190.1 | Unit | 3 | 380 | 20 | 5 |

- Data Analysis Considerations

Because we only include eleven acylcarnitine standards and eleven internal standards in our calibrator curve, the majority of acylcarnitine species are quantitated using a surrogate curve. The table below depicts which acylcarnitines are quantitated using each curve:

| Acylcarnitine Standard/Int. Std. | AC(s) Quantitated Using this Curve |
| --- | --- |
| C2/C2-D3 | C2 |
| C3/C3-D3 | All “C3” species |
| C4/C4-D3 | All “C4” species |
| C5-Iso/C5-Iso-D9 | All “C5” species |
| C6/C6-D3 | All “C6” species |
| C8/C8-D3 | All “C8” species |
| C10/C10-D3 | All “C10” species |
| C12/C12-D3 | All “C12” species |
| C14/C14-D3 | All “C14” species |
| C16/C16-D3 | All “C16” species |
| C18/C18-D3 | All “C18”, “C20” and “C22” species |