General Methods for untargeted metabolomics study conducted in Cornell Proteomics and Metabolomics Facility (for the data generated before Feb2020)

Materials and Methods

Cow Plasma Metabolomics

Sample preparation:

Plasma samples were thawed on ice over a period of 20 minutes. Once thawed, samples were vortexed for approximately 1 minute. Samples (200 L) were transferred to 1.5 mL Eppendorf tubes, and 600 L MeOH (1:3) was added to each plasma sample (refs#1-3). Meanwhile a global QC containing a pool of all the samples of the batch by combining the same volume of each sample) was prepared. Samples were centrifuged, and the supernatant was dried. The samples were re-suspended with 50% acetonitrile (ACN) at 0°C, centrifuged and dried down for final reconstitution in 100 μL of ACN/H2O/Formic Acid (50/50/0.1) containing a std mix (Sulfadimethoxine, 13CPyruvic acid, 13CValine) for liquid chromatography-mass spectrometry (LC-MS) analysis.

LC-MS analysis:

Chromatographic separation was performed on a Vanquish UHPLC system coupled to a Q Exactive™ HF Mass Spectrometer (Thermo Fisher Scientific, San Jose, CA, USA) (Refs#1-3). For HILIC chromatography, a SeQuant ZIC pHILIC column (5μm, 2.1 x 150 mm) was used. The mobile phase consisted of (A) 10 mM ammonium acetate in Water pH=9.8 and (B) acetonitrile. The gradient was as follows: 0–15 min, 90-30% solvent B; 15-18 min, isocratic 30% solvent B; 18-19 min, 30–90% solvent B; 19-27 min, 90% solvent B; followed by 3 min of re-equilibration of the column before the next run. The flow rate was 250 μL/min and the injection volumes were set to 2 μL. To avoid possible bias, the sequence of injections was randomized.

For C18 reverse phase LC, an Accucore Vanquish C18+, 1.5μm column (2.1 mm id x 100mm) was used. The mobile phase consisted of (A) Water /0.1% Formic acid and (B) acetonitrile/0.1% Formic acid. The gradient was as follows: 0 –2.0 min (0.5 - 1%B), 2.0–6.5 min(1–20% B), 6.5-11.5 min (20-95%B), 11.5-13.5 min (95-99% B), 13.5-16.5min (99-100% B), 16.5-19 min (100-0.5% B), 19-24 min(0.5%B). The flow rate was 320  μL/min  and the Injection volumes were set to 2 μL. To avoid possible bias, the sequence of injections was randomized.

All of the samples were analysed by negative and positive electrospray ionization (ESI) in full scan MS mode. Nitrogen as sheath, auxiliary, and sweep gas was set at 50, 8, and 1 U, respectively. Data were acquired under resolving power of 120,000 (at m/z 200); automatic gain control target, 3e6 ions; maximum injection time, 100 ms; scan range, 67-1005 m/z; spray voltage, 3.50 kV; and capillary temperature, 275°C. For QC sample, MS1 data were acquired after every 5-10 samples and used for normalization in quantitation data analysis. ESI −/+ data-dependent MS/MS spectra were generated on the QC sample and used for identification of metabolites. MS/MS data were acquired with a full scan followed by top 15 MS/MS scans with resolving power of 15,000 (at m/z 200); automatic gain control target, 1e5 ions; maximum injection time, 50 ms; isolation window, 0.4 m/z; and stepped NCE 20, 30, and 40.

Data analysis:

The acquired raw files were processed using Compound Discoverer 2.1, an untargeted metabolomics workflow with putative annotation through in-house mass list and databases including ChemSpider, Human Metabolome database (HMDB), Metlin, Kyoto Encyclopedia of Genes and Geneomes (KEGG) and mzCloud were used for processing the raw data and for compound annotation. The software parameters for alignment were 5ppm mass tolerance for the adaptive curve model and 0.5 min maximum shift allowed. For detecting unknown compounds 5 ppm mass tolerance for detection were used along with 30% intensity tolerance, 3 for the signal to noise threshold, and 2 x 106 as minimum peak intensity.

Metabolomics Data Filtering

The resulting data yielded LC-MS features from both positive and negative modes. The process of filtering these features started with removing all features that could not be annotated by Compound Discover 2.1. Removing all those listed as “similar to,” which had a molecular mass that could not be accounted for with the removal or addition of an oxygen or hydrogen group, further narrowed the LC-MS features. Finally, all replicate samples were evaluated. Any features in replicate samples that had significantly different retention times (>1 min) were removed. In some cases, there appeared to be two groups based on retention time within one metabolite. In this case, the known polarity of the metabolite was compared to the retention time. When metabolites were detected in both negative and positive ion modes, those with higher intensity and higher S/N were selected.

Acknowledgements:

We thank the Proteomic and Metabolomics Facility of Cornell University for providing the mass spectrometry data and associated results.

References

- Ledet MM, Harman RM, Fan JC, Schmitt-Matzen E, Diaz-Rubio ME, Zhang S, Van de Walle GR. Secreted sphingomyelins modulate low mammary cancer incidence observed in certain mammals. 2020 Sci Rep. 10(1):20580. doi: 10.1038/s41598-020-77639-1. PMID: 33239740

- Noordwijk KJ, Qin R, Diaz-Rubio ME, Zhang S, Su J, Mahal LK, Reesink HL. Metabolism and global protein glycosylation are differentially expressed in healthy and osteoarthritic equine carpal synovial fluid. Equine Vet J. 2021 Feb 15;. doi: 10.1111/evj.13440. [Epub ahead of print] PubMed PMID: 33587757.

- Bhawal R, Fu Q, Anderson ET, Gibson GE, Zhang S. Serum Metabolomic and Lipidomic Profiling Reveals Novel Biomarkers of Efficacy for Benfotiamine in Alzheimer's Disease Int J Mol Sci. 2021 22(24):13188. doi: 10.3390/ijms222413188. PMID: 34947984.