Separation was performed by Ultimate 3000LC combined with Q Exactive 625 MS (Thermo) and screened with ESI-MS (targeted MS/MS mode). Mass spectrometry parameters in ESI+ and ESI- mode are listed as follows: 632 ESI+: Heater Temp 300 °C; Sheath Gas Flow rate, 45 arb; Aux Gas Flow Rate, 15 arb; Sweep Gas 633 Flow Rate, 1 arb; spray voltage, 3.0 kV; Capillary Temp, 350 °C; S-Lens RF Level, 30%. 634 ESI-: Heater Temp 300 °C, Sheath Gas Flow rate, 45 arb; Aux Gas Flow Rate, 15arb; Sweep Gas 635 Flow Rate, 1 arb; spray voltage, 3.2 kV; Capillary Temp,350 °C; S-Lens RF Level, 60%. The software we used for metabolites identification is Compound Discoverer 3.1 SP1. For metabolites identification, the raw data are acquired and aligned using the Compound Discover based on the m/z value and the retention time of the ion signals. CD software offers a fully integrated suite of advanced software tools for known-parent and unknown data processing and interpretation. It is used for the initial metabolites identification. We also use mzMine2 software (with parameters of m/z tolerance=0.1 m/z ) for the online database search based on the MS1 and MS2 spectra acquired. The database used is HMDB.