| Malaria Host-Pathogen Interaction Center - MaHPIC Systems Biology of Malaria as a Model for Host/Pathogen Interactions Contract HHSN272201200031C |  |
| --- | --- |
| Institution: | Emory University |
| Core | Metabolomics |
| SOP | Orbitrap Setup |
| Date | May 22 2013 |

## Distribution

Esmeralda Meyer
MaHPIC Project Manager

## Revision Record

| Revision | Date | Responsible Person | Description of change |
| --- | --- | --- | --- |
| 1 | May 17 2013 | ViLinh Tran | Initial Release |

Orbitrap Instrument Setup

Purpose: Instructions on the preparation before starting a batch of samples and how to operate the instrument to start the batch

Material:

Table 1. Reagents for LC solvents

| Product Name | Vendor | Part Number |
| --- | --- | --- |
| Acetonitrile HPLC 99.9% | Sigma-Aldrich | 34998-4L |
| Methanol HPLC 99.9% | Sigma-Aldrich | 34860-4L-R |
| LCMS water | Fluka | 39253-4L-R |
| Formic acid 98% | Fluka | 94318-250ML-F |

Procedures:

With sample preparation done and loaded onto the autosampler plates; the following parameters have to be met before starting a batch

- Check solvent bottles (found behind the LC instrument)

  - Both acetonitrile bottles need daily refills; water is refilled every 2-3 days (bottles are labeled)

    - Solvent A: LCMS water with 2% Formic Acid

    - Solvent B: LCMS Acetonitrile

    - Solvent C: LCMS water

    - Solvent D: LCMS Acetonitrile with 2% Formic Acid

  - Dynamic Load and Wash (DLW) solvents: 50:50 MeOH: H2O (blue lid bottles)

    - Refilled once every 3 weeks (or when less than 300mL left in bottle)

  - Solvent A and D: Two separate graduated cylinder with its respective label

    - Calculate 2% of the total volume

      - Ex. 2% of 500mL is 10mL.

    - Pipette 10mL of Formic acid into the graduated cylinder and fill the graduated cylinder up to 500mL.

    - Use funnel to pour the solution into the correct solvent bottles

- Prepare a sequence file for each batch (can be prepared on your own computer on EXCEL)

  - Export a template from Xcalibur or use an existing template located at G:\Pulmonary\Research\Jones Lab\Orbitrap\Sequence Files\Sequence File Templates

    - The templates are made for each rack (1-6) and also based on a triplicate run (20 samples per batch)

  - For each study, give each batch a separate sequence file and name as following:

    - Technician Initials_Studyname_batch[number]_date.csv

      - VT_Marmoset_batch1_101811.csv

      - VT_Marmoset_batch2_101911.csv

Orbitrap Instrument Setup

  - SAVE sequence file as .csv file and save it to G:\Pulmonary\Research\Jones Lab\Orbitrap\Sequence Files\[folder with study name]

    - Will be imported into Xcalibur (File > import files > choose correct .csv)

  - Each template consists of the same columns:

    - Sample type, File Name, Sample ID, Path, Instrument Method, Position, Inj Vol (see Table 6)

      - Sample Type – always UNKNOWN

      - File Name – name of samples given by the PI of the study

      - Sample ID – Orbitrap ID (see Procedures 2b)

      - Path – output files of RAW data

      - Instrument Method: Prepare gradient for each column (see Table 4 and 5)

      - Position – location of the samples on rack

      - Injection Volume – 10uL per run

| Table 2. Mobile phase gradient for anion exchange (AE) analysis |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Time (min) | Solvent A (2% Formic Acid in H2O), % | Solvent B (ACN), % | Solvent C (H2O), % | Solvent D (2% Formic Acid in ACN), % | Flow rate (L/min) |
| 0 | 5 | 50 | 45 | 0 | 350 |
| 2 | 5 | 50 | 45 | 0 | 350 |
| 7 | 50 | 50 | 0 | 0 | 350 |
| 10 | 50 | 50 | 0 | 0 | 350 |

| Table 3. Mobile phase gradient for C18 analysis |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Time (min) | Solvent A (2% Formic Acid in H2O), % | Solvent B (ACN), % | Solvent C (H2O), % | Solvent D (2% Formic Acid in ACN), % | Flow rate (L/min) |
| 0 | 5 | 35 | 60 | 0 | 350 |
| 2 | 5 | 35 | 60 | 0 | 350 |
| 7 | 5 | 95 | 0 | 0 | 400 |
| 10 | 5 | 95 | 0 | 0 | 400 |

  - Only the first sequence should contain the following: Equilibrate, AE std, C18 std, washin, nist, and qstd (other sets will not include equilibrate, ae, c18, and wash-in)

Table 4. Example of template for the sequence file

| Sample Type | Sample ID | File Name | Path | Instrument Method | Position | Inj Vol |
| --- | --- | --- | --- | --- | --- | --- |
| Unknown | Equil | Equilibrate | C:\Xcalibur\Data | C:\Xcalibur\methods\Equilibrate | CStk1-01:1 | 10 |
| Unknown | Ae std | Ae std | C:\Xcalibur\Data | C:\Xcalibur\methods\AnionExchange | CStk1-01:2 | 10 |
| Unknown | C18 std | C18 std | C:\Xcalibur\Data | C:\Xcalibur\methods\c18 | CStk1-01:3 | 10 |
| Unknown | Washin | Washin | C:\Xcalibur\Data | C:\Xcalibur\methods\AnionExchange | CStk1-01:4 | 10 |
| Unknown | Washin | Washin | C:\Xcalibur\Data | C:\Xcalibur\methods\c18 | CStk1-01:4 | 10 |
| Unknown | Washin | Washin | C:\Xcalibur\Data | C:\Xcalibur\methods\AnionExchange | CStk1-01:4 | 10 |
| Unknown | Washin | Washin | C:\Xcalibur\Data | C:\Xcalibur\methods\c18 | CStk1-01:4 | 10 |
| Unknown | Washin | Washin | C:\Xcalibur\Data | C:\Xcalibur\methods\AnionExchange | CStk1-01:4 | 10 |
| Unknown | Washin | Washin | C:\Xcalibur\Data | C:\Xcalibur\methods\c18 | CStk1-01:4 | 10 |
| Unknown | Washin | Washin | C:\Xcalibur\Data | C:\Xcalibur\methods\AnionExchange | CStk1-01:4 | 10 |
| Unknown | Washin | Washin | C:\Xcalibur\Data | C:\Xcalibur\methods\c18 | CStk1-01:4 | 10 |
| Unknown | Washin | Washin | C:\Xcalibur\Data | C:\Xcalibur\methods\AnionExchange | CStk1-01:4 | 10 |
| Unknown | Washin | Washin | C:\Xcalibur\Data | C:\Xcalibur\methods\c18 | CStk1-01:4 | 10 |
| Unknown | Nist1-1 | VT_111006_001 | C:\Xcalibur\Data | C:\Xcalibur\methods\AnionExchange | CStk1-01:5 | 10 |
| Unknown | Nist1-2 | VT_111006_002 | C:\Xcalibur\Data | C:\Xcalibur\methods\c18 | CStk1-01:5 | 10 |
| Unknown | Nist1-3 | VT_111006_003 | C:\Xcalibur\Data | C:\Xcalibur\methods\AnionExchange | CStk1-01:5 | 10 |
| Unknown | Nist1-4 | VT_111006_004 | C:\Xcalibur\Data | C:\Xcalibur\methods\c18 | CStk1-01:5 | 10 |
| Unknown | Qstd1-1 | VT_111006_005 | C:\Xcalibur\Data | C:\Xcalibur\methods\AnionExchange | CStk1-01:6 | 10 |
| Unknown | Qstd1-2 | VT_111006_006 | C:\Xcalibur\Data | C:\Xcalibur\methods\c18 | CStk1-01:6 | 10 |
| Unknown | Qstd1-3 | VT_111006_007 | C:\Xcalibur\Data | C:\Xcalibur\methods\AnionExchange | CStk1-01:6 | 10 |
| Unknown | Qstd1-4 | VT_111006_008 | C:\Xcalibur\Data | C:\Xcalibur\methods\c18 | CStk1-01:6 | 10 |

- Make sure the positions are labeled correctly

  - 3 shelves with 2 racks per shelf

  - Label correctly: ex. CStk1-01:1

    - CStk1 will always be constant and will not change

    - 01 represents the number of the rack

      - 1, 3, and 5 are the racks towards the front of the instrument

      - 2,4, and 6 are the racks towards the back of the instrument

    - The last number represents the sample number on the rack

  - Racks 1 and 2 belongs to the Top shelf; Racks 3 & 4 are located on the middle shelf; racks 5 & 6 are located on the bottom shelf

- Optimized settings of the instrument

  - HESI probe with S-lens combination for ESI

  - MS1 mode scanning m/z range of 85-2000

  - Resolution - 60,000

  - maximum number of ions collected - 5.00 x 105

  - The maximum injection time - 5 L/s

  - capillary temperature - 275 oC

  - source heater - 45 oC

  - voltage - 4.6 kV

  - sheath gas – 45flow rate

  - auxillary gas flow – 5

  - sweep gas flow – 0

  - Each sample is run in triplicate with dual chromatography (AE and C18) and 10 μL injection volume on each column.

    - An example of analysis with apLCMS [4] showed 13,034 m/z features using AE.  Analysis on C18 resolved 14,317 m/z features.

- Once everything is checked out and refilled

  - Highlight the files you want to run and then click on the  icon to queue the files

  - If anything were to happen and you need to stop the current run

    - Click on the  icon. This will continue until it finishes the current run and pause before the next.

      - Once the pause icon is clicked again – the run will continue onto the next run. *caution – always double check to make sure*

    - You can always hit the  icon immediately after pressing the pause to stop all the runs completely

      - If pause and then stop – the queue will be completely cleared and the samples will need to be highlighted and again and    clicked again

Appendix A-1

Important Contact Information

Purpose: Provide contact information for assistance of any malfunctions to the LTQ Velos Orbitrap and the CTC PAL autosampler

| Name | Company | Title | Phone | Email address |
| --- | --- | --- | --- | --- |
| Service call | Thermo Fisher | Tech Support | 800-532-4752 | Click on 1 to open service call |
| Steven Fannin | Thermo Fisher | Sales Representative | 678-957-8498 | Steve.fannin@thermofisher.com |
| Mark Mager | Thermo Fisher | Engineer | 850-206-1629 | Mark.mager@thermofisher.com |
| Jeff Cheeves | Thermo Fisher | Sales, consumables | 404-732-4422 | Jeff.cheeves@thermofisher.com |
| Alisha Roberts | Thermo Fisher | Tech Support | 561-687-6206 | Alisha.roberts@thermofisher.com |
| Bill Mahn | Thermo Fisher | Tech Support |  | Bill.mahn@thermofisher.com |
| Troy Gorbett | Thermo Fisher | CDC orbitrap engineer | 404-502-0921 | Troy.gorbett@thermofisher.com |
| Scott Davidson | Thermo Fisher | Accela engineer | 904-540-8218 | Scott.davidson@thermofisher.com |
| Mark Netsch | Thermo Fisher | Accela engineer | 508-963-9362 | Mark.netsch@thermofisher.com |
| Andy Tai | Thermo Fisher | Product support engineer |  | Andy.tai@thermofisher.com |
| Sunny Qiu | Thermo Fisher | CDC Engineer | 404-368-0395 | Sunny.qiu@thermofisher.com |
| Michael Burns | Peak Scientific | Sales Manager (n2 generator) | 8666471649 | mburns@peakscientific.com |
| Alison McEwan | Peak Scientific | Sales Manger (n2 generator) | 858-822-8894 | amcwean@peakscientific.com |
| Jarrett Cooper | Thermo Fisher | Tech support | 8005324752 ext 3696257 | jarrett.cooper@thermofisher.com |
| Steven Jones | Higgins Analytical | General Manager | 650-988-8930 | slj@higginsanalytical.com |

Appendix A-2

Important Contact Information

Procedures

- Instrument maintenance and basic troubleshooting will be discussed further in another chapter of the manual

- If after troubleshooting the problem has not been fixed; prepare the following before contacting tech support

  - Pictures of any error messages or the area on the instrument that is malfunctioning

  - Prepare a detailed description of the problem and provide all the troubleshooting details that have already been done

  - Have the serial number of the instrument ready (Appendix B-1)

  - Version numbers of the programs and instrument

    - This will give the tech support a better idea of how old the system is and which manual to use (Appendix C-1)

- Call the following for specific problems

  - LTQ Velos Orbitrap – Troy Gorbett

  - CTC PAL autosampler – Scott Davidson

  - Miscellaneous/Xcalibur – Alicia Roberts

- Tech/engineer will usually place the service call after discussion of the problems, however, you can put in a service call yourself by dialing Thermo Fisher Tech support line at 1800-532-4752, and pressing 1 to open a service call

Appendix B-1

Serial Numbers

| LTQ Orbitrap Velos | SN 03443B |
| --- | --- |
| PAL HTC Accela1-TMO (autosampler) | SN 211980 |
| Pump 1 unit | SN 25373 |
| Pump 2 unit | SN 25555 |
| N2 generator Model: NM32LA | SN A13-02-001 |

Appendix C-1

Version Information

Foundation	Version 1.0.0.42

- Applications

  - FT programs 2.1.0.40

  - Xcalibur 2.1.0.1140

- Configured instruments

  - Accela Open AS 2.5.0

  - Accela 1250 pump 1.05.0900

  - LTQ Orbitrap Velos MS 2.6.0.1050

Appendix D-1

Materials

| Product name | Vendor | Part number |
| --- | --- | --- |
| 1.5mL Microcentrifuge Natural Tube | USA Scientific | 1415-2500 |
| 11mm pp vial crimp/snap 250uL | National Scientific | C4011-13 |
| 11mm pink snap cap, Red T/S 100/pk | Thermo Scientific | 60180-677 |
| 0.1-10uL Natural Pipette Tips | Tip One | 1111-3700 |
| 1-200uL Yellow Pipette Tips | Tip One | 1111-1706 |
| 1000uL Blue Pipette Tips | Tip One | 1111-2721 |

Reagents

| Product Name | Vendor | Part Number |
| --- | --- | --- |
| Acetonitrile HPLC 99.9% | Sigma-Aldrich | 34998-4L |
| Methanol HPLC 99.9% | Sigma-Aldrich | 34860-4L-R |
| LCMS water | Fluka | 39253-4L-R |
| Formic acid 98% | Fluka | 94318-250ML-F |

Calibration Standards

| Standards | Vendor | Part number |
| --- | --- | --- |
| Caffeine | Sigma | C1778 |
| MRFA | Research Plus, Inc. | 18001 |
| Ultramark 1621 | Alfa Aesar | L16698 |
| N-butylamine | Sigma | 471305-25mL |
| Sodium dodecyl sulfate | Sigma | L4509-10g |
| Sodium taurocholate | Sigma | T4009-250mg |

Columns

| Type of column | Vendor | Part Number |
| --- | --- | --- |
| Anion exchange 100x2.1mm,7um | Hamilton | PRP-X110S |
| C18 100x2.1mm, 5um | Higgins | TS-1021-C185 |
| Guard column c18 20x3.2mm,5um | Higgins | TF-0232-C185 |