Newman Lab Lipid Mediator Analysis

Materials & Methods

For plasma analysis, oxylipins, endocannabinoids, and ceramides were isolated using a Waters Ostro Sample Preparation Plate (Milford, MA). Sample aliquots (50 μL plasma) were introduced into the plate wells and spiked with a 5 µL anti-oxidant solution (0.2 mg/ml solution BHT/EDTA in 1:1 MeOH:water) and 10 μL 1000nM analytical deuterated surrogates. Acetonitrile (150 μL) with 1% formic acid was forcefully added to the sample and aspirated three times to mix. Samples were eluted into glass inserts containing 10 μL 20% glycerol by applying a vacuum at 15 Hg for 10 min. Eluent was dried by speed vacuum for 35 min at the medium BP setting, before switching to an aqueous setting for an additional 35 min. Once dry, samples were re-constituted with the internal standard 1-cyclohexyl ureido, 3-dodecanoic acid (CUDA) and 1-Phenyl 3-Hexadecanoic Acid Urea (PHAU) at 100 nM (50:50 MeOH:CAN), vortexed 1 min, transferred to a spin filter (0.1 µm, Millipore, Billerica, MA), centrifuged for 3 min at 6ºC at <4500g (rcf), before being transferred to 2 mL LC-MS amber vials. Extracts were stored at -20ºC until analysis by UPLC-MS/MS. The internal standard was used to quantify the recovery of surrogate standards.

For tissue analysis (liver, muscle, adipose, and hypothalamus), samples were pulverized, added to 2mL polypropylene tubes, and spiked with a 5 µL anti-oxidant solution (0.2 mg/ml solution BHT/EDTA in 1:1 MeOH:water) and 10 μL 1000nM analytical deuterated surrogates. A total of 50 µL of methanol was added and the tube was placed in a Geno/Grinder for 30 sec. An additional 550µL isopropanol w/ 10mM ammonium formate & 1% formic acid and 100 uL water was added and the tube was placed in a Geno/Grinder for 30 sec before being centrifuged at 10,000g for 5 min at room temp. The supernate was then transferred into Water Ostro Sample Preparation Plate wells. From this point forward, the samples were treated similarly to plasma (detailed previously above), as they were eluted by vacuum, dried, reconstituted, and analyzed.

Analytes in a 100 μL extract aliquots were separated utilizing a Waters Acquity UPLC (Waters, Milford, MA) with a solvent gradient using modifications of a previously published protocols for oxylipins (1), endocannabinoids (2), and ceramides (3)(See pages 3-11 for detailed instrument parameters for each assay). Samples were held at 10ºC. Separated residues were detected by negative mode electrospray ionization for oxylipins and positive mode electrospray ionization for endocannabinoids and ceramides using multiple reaction monitoring on an API 4000 QTrap (AB Sciex, Framingham, MA, USA). Analytes were quantified using internal standard methods and 5 to 10 point calibration curves (r2 ≥ 0.997). Calibrants and internal standards were either synthesized [10,11-DHN, 10,11-DHHep, 10(11)-EpHep and CUDA] or purchased from Cayman Chemical (Ann Arbor, MI), Avanti Polar Lipids Inc. (Alabaster, AL), and Sigma-Aldrich (St. Louis, MO), unless otherwise indicated. Loradan Fine Lipids (Malmo, Sweden) provided the linoleate derived triols 9,12,13-TriHOME and 9,10,13- TriHOME. Data was processed utilizing AB Sciex MultiQuant version 3.0.1.

- Strassburg K et al (2012). Quantitative profiling of oxylipins through comprehensive LC-MS/MS analysis: application in cardiac surgery. Anal Bioanal Chem. 404:1413-26.

- Grapov D et al (2012). Type 2 Diabetes Associated Changes in the Plasma Non-Esterified Fatty Acids, Oxylipins and Endocannabinoids. PLoS ONE. 7(11):e48852.

- Bielawski J et al (2009). Comprehensive quantitative analysis of bioactive sphingolipids by high-performance liquid chromatography-tandem mass spectrometry. Methods Mol Biol. 579:443-67.

| Table 1a. Oxylipin Assay UPLC parameters |  |  |
| --- | --- | --- |
| Time (min) | A% | B% |
| 0 | 90 | 70 |
| 2 | 90 | 80 |
| 5 | 60 | 85 |
| 5.5 | 60 | 90 |
| 13.5 | 20 | 95 |
| 13.75 | 20 | 99 |
| 14.5 | 5 | 99 |
| 14.7 | 5 | 70 |
| 15.2 | 90 | 70 |
| Solvent A = 5 mm NH4COO 0.2% |  |  |
| formic acid; Solvent B = 5 mm |  |  |
| NH4COO 0.2% formic acid in MeOH, |  |  |
| flow rate = 0.25 mL/min, column |  |  |
| 2.1 X 100mm, 1.7 µm BEH C8 (Waters, |  |  |
| Milford, MA), column temp = 60 °C |  |  |

| Table 1b. Endocannabinoid Assay UPLC parameters |  |  |  |
| --- | --- | --- | --- |
| Time (min) | A% | B% |  |
| 0 | 75 | 25 |  |
| 0.25 | 75 | 25 |  |
| 0.5 | 60 | 40 |  |
| 1.5 | 50 | 50 |  |
| 3 | 45 | 55 |  |
| 3.5 | 20 | 80 |  |
| 8 | 15 | 85 |  |
| 9 | 5 | 95 |  |
| 9.25 | 5 | 95 |  |
| 9.5 | 75 | 25 |  |
| 11 | 75 | 25 |  |
| Solvent A = 0.1% Acetic Acid; |  |  |  |
| Solvent B = 90% Acetonitrile / 10% |  |  |  |
| isopropanol flow rate = 0.25 mL/min, |  |  |  |
| column 2.1 X 150mm, 1.7 µm BEH C18 |  |  |  |
| (Waters, Milford, MA), column |  |  |  |
| temp = 60 °C |  |  |  |

| Table 1c. Ceramide Assay UPLC parameters |  |  |
| --- | --- | --- |
| Time (min) | A% | B% |
| 0 | 90 | 70 |
| 2 | 90 | 80 |
| 5 | 60 | 85 |
| 5.5 | 60 | 90 |
| 13.5 | 20 | 95 |
| 13.75 | 20 | 99 |
| 14.5 | 5 | 99 |
| 14.7 | 5 | 70 |
| 15.2 | 90 | 70 |
| Solvent A = 5 mm NH4COO 0.2% |  |  |
| formic acid; Solvent B = 5 mm |  |  |
| NH4COO 0.2% formic acid in MeOH, |  |  |
| flow rate = 0.25 mL/min, column |  |  |
| 2.1 X 100mm, 1.7 µm BEH C8 (Waters, |  |  |
| Milford, MA), column temp = 60 °C |  |  |

|  | Table 2a. Oxylipin UPLC/ESI QTRAP Analyte and Instrument-specific Parameters * † †† |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Name | Expected RT | Internal Standard (ISTD) Name | Parent ion | Daughter ion |  |
| 1 | PHAU | 2.67 |  | 249.2 | 130.1 |  |
| 2 | 6-keto PGF1a | 3.04 | IS PGD2 | 369.3 | 163.1 |  |
| 3 | d4-6-Keto PGF1a | 3.04 | IS PHAU | 373.3 | 167.1 |  |
| 4 | PGF3a | 3.56 | IS d4-PGF2a | 351.3 | 307.4 |  |
| 5 | PGE3 | 3.69 | IS d4-PGF2a | 349.3 | 269.2 |  |
| 6 | d4-TXB2 | 3.74 | IS PHAU | 373.3 | 173.1 |  |
| 7 | TXB2 | 3.75 | IS d4-TXB2 | 369.3 | 169.1 |  |
| 8 | 9_12_13-TriHOME | 4.22 | IS d4-PGF2a | 329.2 | 211.2 |  |
| 9 | d4-PGF2a | 4.27 | IS PHAU | 357.3 | 197.2 |  |
| 10 | PGF2a - Select | 4.28 | IS d4-PGF2a | 353.3 | 193.2 |  |
| 11 | PGE2 | 4.40 | IS d4-PGD2 | 351.3 | 271.2 |  |
| 12 | PGE1 | 4.58 | IS d4-PGD2 | 353.3 | 317.2 |  |
| 13 | d4-PGD2 | 4.69 | IS PHAU | 355.3 | 275.2 |  |
| 14 | PGD2 | 4.71 | IS d4-PGD2 | 351.3 | 271.2 |  |
| 15 | 15-Keto PGE2 | 4.75 | IS d4-PGD2 | 349.2 | 331.3 |  |
| 16 | Resolvin D1 | 5.23 | IS d4-PGF2a | 375.3 | 121.1 |  |
| 17 | Lipoxin A4 | 5.35 | IS d4-PGF2a | 351.3 | 217.2 |  |
| 18 | LTB5 | 6.40 | IS d4-LTB4 | 333.3 | 195.2 |  |
| 19 | d3-LTE4 | 6.51 | IS CUDA | 441.4 | 336.3 |  |
| 20 | LTE4 | 6.54 | IS d3-LTE4 | 438.4 | 333.3 |  |
| 21 | 15_16-DiHODE | 6.85 | IS d11-14_15-DiHETrE | 311.2 | 235.2 |  |
| 22 | 12_13-DiHODE | 6.93 | IS d11-14_15-DiHETrE | 311.2 | 183.2 |  |
| 23 | 8_15-DiHETE | 6.95 | IS d11-14_15-DiHETrE | 335.3 | 235.2 |  |
| 24 | 9_10-DiHODE | 6.98 | IS d11-14_15-DiHETrE | 311.2 | 201.2 |  |
| 25 | 17_18-DiHETE | 7.23 | IS d11-14_15-DiHETrE | 335.3 | 247.2 |  |
| 26 | 5_15-DiHETE | 7.28 | IS d11-14_15-DiHETrE | 335.3 | 173.1 |  |
| 27 | 6-trans-LTB4 | 7.40 | IS d4-LTB4 | 335.3 | 195.2 |  |
| 28 | 14_15-DiHETE | 7.57 | IS d11-14_15-DiHETrE | 335.3 | 207.2 |  |
| 29 | CUDA | 7.62 |  | 339.4 | 214.2 |  |
| 30 | d4-LTB4 | 7.63 | IS CUDA | 339.3 | 163.1 |  |
| 31 | LTB4 | 7.67 | IS d4-LTB4 | 335.3 | 195.2 |  |
| 32 | 12_13-DiHOME | 7.94 | IS d11-14_15-DiHETrE | 313.3 | 183.2 |  |
| 33 | 10_11-DHHep | 8.02 | IS CUDA | 301.2 | 283.2 |  |
| 34 | 9_10-DiHOME | 8.35 | IS d11-14_15-DiHETrE | 313.3 | 201.2 |  |
| 35 | d11-14_15-DiHETrE | 8.51 | IS CUDA | 348.4 | 207.2 |  |
| 36 | 19_20-DiHDoPA | 8.57 | IS d11-14_15-DiHETrE | 361.3 | 273.20 |  |
| 37 | 14_15-DiHETrE | 8.60 | IS d11-14_15-DiHETrE | 337.3 | 207.2 |  |
| 38 | 11_12-DiHETrE | 9.16 | IS d11-14_15-DiHETrE | 337.3 | 167.1 |  |
| 39 | 9_10-e-DiHO | 9.31 | IS 10_11-DHHep | 315.2 | 297.2 |  |
| 40 | 9-HOTE | 9.42 | IS d4-9(S)-HODE | 293.2 | 171.1 |  |
| 41 | 12(13)-Ep-9-KODE | 9.43 | IS d4-9(S)-HODE | 309.2 | 291.2 |  |
| 42 | 13-HOTE | 9.55 | IS d4-9(S)-HODE | 293.2 | 195.2 |  |
| 43 | 8_9-DiHETrE | 9.70 | IS d11-14_15-DiHETrE | 337.3 | 127.1 |  |
| 44 | 15-deoxy PGJ2 | 9.77 | IS d11-14_15-DiHETrE | 315.2 | 271.2 |  |
| 45 | d6-20-HETE | 9.93 | IS CUDA | 325.3 | 281.2 |  |
| 46 | 15-HEPE | 9.97 | IS d8-12(S)-HETE | 317.2 | 219.2 |  |
| 47 | 20-HETE | 9.97 | IS d6-20-HETE | 319.2 | 275.2 |  |
| 48 | 12-HEPE | 10.33 | IS d8-12(S)-HETE | 317.2 | 179.1 |  |
| 49 | 5_6-DiHETrE | 10.46 | IS d11-14_15-DiHETrE | 337.3 | 145.1 |  |
| 50 | 9-HEPE | 10.53 | IS d4-9(S)-HODE | 317.2 | 167.2 |  |
| 51 | 13-HODE | 10.83 | IS d4-9(S)-HODE | 295.2 | 195.2 |  |
| 52 | 5-HEPE | 10.91 | IS d4-9(S)-HODE | 317.2 | 115.1 |  |
| 53 | d4-9(S)-HODE | 10.95 | IS CUDA | 299.2 | 172.1 |  |
| 54 | 9-HODE | 11.01 | IS d4-9(S)-HODE | 295.2 | 171.1 |  |
| 55 | 15(16)-EpODE | 11.09 | IS d4-12(13)-EpOME | 293.2 | 275.2 |  |
| 56 | 17(18)-EpETE | 11.19 | IS d4-12(13)-EpOME | 317.2 | 259.2 |  |
| 57 | 15-HETE | 11.21 | IS d8-12(S)-HETE | 319.2 | 219.2 |  |
| 58 | 13-KODE | 11.27 | IS d4-9(S)-HODE | 293.2 | 179.1 |  |
| 59 | 9(10)-EpODE | 11.28 | IS d4-12(13)-EpOME | 293.2 | 275.2 |  |
| 60 | 17-HDoHE | 11.32 | IS d8-12(S)-HETE | 343.3 | 281.2 |  |
| 61 | 15-HpETE screen | 11.48 | IS d8-12(S)-HETE | 335.2 | 113.1 |  |
| 62 | 13-HpODE screen | 11.48 | IS d4-9(S)-HODE | 311.2 | 179.1 |  |
| 63 | 9-HpODE screen | 11.88 | IS d4-9(S)-HODE | 311.2 | 185.2 |  |
| 64 | 12(13)-EpODE | 11.48 | IS d4-12(13)-EpOME | 293.2 | 183.2 |  |
| 65 | 15-KETE | 11.55 | IS d8-12(S)-HETE | 317.2 | 273.2 |  |
| 66 | 14-HDoHE | 11.62 | IS d8-12(S)-HETE | 343.3 | 281.2 |  |
| 67 | 11-HETE | 11.64 | IS d8-12(S)-HETE | 319.2 | 167.1 |  |
| 68 | 14(15)-EpETE | 11.64 | IS d4-12(13)-EpOME | 317.2 | 247.2 |  |
| 69 | 9-KODE | 11.71 | IS d4-9(S)-HODE | 293.2 | 185.2 |  |
| 70 | d8-12(S)-HETE | 11.78 | IS CUDA | 327.2 | 184.2 |  |
| 71 | 11(12)-EpETE | 11.80 | IS d4-12(13)-EpOME | 317.2 | 167.3 |  |
| 72 | 12-HETE | 11.88 | IS d8-12(S)-HETE | 319.2 | 179.1 |  |
| 73 | 8-HETE | 12.02 | IS d8-12(S)-HETE | 319.2 | 155.1 |  |
| 74 | 9-HETE | 12.23 | IS d8-12(S)-HETE | 319.2 | 167.1 |  |
| 75 | 12-HpETE screen | 12.00 | IS d8-12(S)-HETE | 335.2 | 153.1 |  |
| 76 | 5-HpETE screen | 13.18 | IS d8-5(S)-HETE | 335.2 | 155.1 |  |
| 77 | d8-5(S)-HETE | 12.49 | IS CUDA | 327.2 | 116.1 |  |
| 78 | 19(20)-EpDPE | 12.57 | IS d4-12(13)-EpOME | 343.3 | 281.2 |  |
| 79 | 5-HETE | 12.58 | IS d8-5(S)-HETE | 319.2 | 115.1 |  |
| 80 | d4-12(13)-EpOME | 12.65 | IS CUDA | 299.2 | 198.1 |  |
| 81 | 12(13)-EpOME | 12.74 | IS d4-12(13)-EpOME | 295.2 | 195.2 |  |
| 82 | 14(15)-EpETrE | 12.83 | IS d4-12(13)-EpOME | 319.2 | 219.2 |  |
| 83 | 4-HDoHE | 12.88 | IS d8-5(S)-HETE | 343.3 | 281.2 |  |
| 84 | 16(17)-EpDPE | 12.96 | IS d4-12(13)-EpOME | 343.5 | 273.5 |  |
| 85 | 9(10)-EpOME | 12.97 | IS d4-12(13)-EpOME | 295.2 | 171.1 |  |
| 86 | 5-KETE | 13.28 | IS d8-5(S)-HETE | 317.2 | 203.2 |  |
| 87 | 11(12)-EpETrE | 13.31 | IS d4-12(13)-EpOME | 319.2 | 167.1 |  |
| 88 | 8(9)-EpETrE | 13.51 | IS d4-12(13)-EpOME | 319.2 | 155.1 |  |
| 89 | 10-Nitrolinoleate | 13.79 | IS d17-10-Nitrooleate | 324.3 | 277.2 |  |
| 90 | 9_10-EpO | 14.37 | IS 10_11-DHHep | 297.3 | 279.2 |  |
| 91 | d17-10-Nitrooleate | 14.60 | IS CUDA | 343.2 | 307.5 |  |
| 92 | 10-Nitrooleate | 14.7 | IS d17-10-Nitrooleate | 326.2 | 279.5 |  |
| 93 | 9-Nitrooleate | 14.80 | IS d17-10-Nitrooleate | 326.2 | 308.2 |  |
| 94 | 9_10-DiHHex | 15.21 | IS 10_11-DHHep | 287.2 | 227.5 |  |
| 95 | d8-Arachidonic Acid | 15.68 | IS CUDA | 311.3 | 267.1 |  |
| * - Analytes were separated under conditions described in Table Ia. Collision-induced dissociation was performed with nitrogen |  |  |  |  |  |  |
| at a pressure of 2.3 mTorr. Scheduled mass spectral multiple reaction monitoring was utilized to identify analytes. |  |  |  |  |  |  |
| † - Internal Standards (ISTD) - Analytes were corrected for recoveries of listed surrogates. 1-Cyclohexylureido,3-dodecanoic acid |  |  |  |  |  |  |
| (CUDA) and 1-Phenyl 3-Hexadecanoic Acid Urea (PHAU) were introduced immediately prior to analysis and used to quantify |  |  |  |  |  |  |
| surrogate recoveries. |  |  |  |  |  |  |
| †† - Compounds labeled as "screen" are compounds for which we did not have calibration standards. The |  |  |  |  |  |  |
| compounds were identified based on retention time and transition (Da) and produced qualitative data. |  |  |  |  |  |  |

Table 2b. UPLC/Electrospray ionization QTRAP analyte and instrument specific parameters*

| Analyte | Common Abbreviation | tR (min) | Transition (Da) | Dwell (msec) | Declustering (V) | Collision (V) | ISTD† |
| --- | --- | --- | --- | --- | --- | --- | --- |
| d4-PGF2α Ethanolamide | d4-PGF2α EA | 3.24 | 384.3 > 62.1 | 25 | 55 | 40 | CUDA |
| PGF2αEthanolamide | PGF2α EA | 3.25 | 380.3 > 62.1 | 25 | 55 | 41 | d4-PGF2a EA |
| PGE2 Ethanolamide | PGE2 EA | 3.26 | 378.3 > 62.1 | 25 | 58 | 38 | d4-PGF2a EA |
| PGD2 Ethanolamide | PGD2 EA | 2.9 | 378.3 > 62.1 | 25 | 58 | 38 | d4-PGF2a EA |
| PGF2α 1-glyceryl ester | PGF2a 1G | 3.04 | 411.3 > 301 | 25 | 55 | 19 | d4-PGF2a EA |
| PGE2 1-glyceryl ester | PGE2 1G | 3.07 | 409.3 > 317 | 25 | 55 | 19 | d4-PGF2a EA |
| 1-Cyclohexyluriedo-3-dodecanoic acid | CUDA | 4.98 | 341.3 > 216 | 25 | 58 | 25 | --- |
| 15(S)-HETE Ethanolamide | 15-HETE EA | 5.19 | 346.3 > 62.1 | 25 | 58 | 39 | d8-AEA |
| 11(12)-EET Ethanolamide | 11(12)-EpETre EA | 5.58 | 364.3 > 62.1 | 25 | 58 | 40 | d8-AEA |
| α-Linolenoyl Ethanolamide | αLEA | 6.03 | 322.2 > 62.1 | 25 | 72 | 32 | d8-AEA |
| Docosahexaenoyl Ethanolamide | DHEA | 6.42 | 372.3 > 62.1 | 25 | 61 | 36 | d8-AEA |
| d8-Arachidonoyl Ethanolamide | d8-AEA | 6.54 | 356.3 > 63.1 | 25 | 60 | 30 | CUDA |
| Arachidonoyl Ethanolamide | AEA | 6.59 | 348.3 > 62.1 | 25 | 65 | 33 | d8-AEA |
| Linoleoyl Ethanolamide | LEA | 6.66 | 324.2 > 62.1 | 25 | 72 | 31 | d8-AEA |
| d5-2-Arachidonoyl Glycerol | d5-2-AG | 6.92 | 384.3 > 287 | 20 | 63 | 19 | CUDA |
| 2-Arachidonoyl Glycerol | 2-AG | 6.94 | 379.3 > 287 | 20 | 53 | 19 | d5-2-AG |
| d8-Arachidonoyl Glycine | d8-NA-Gly | 6.95 | 370.3 > 76.1 | 25 | 79 | 35 | CUDA |
| Arachidonoyl Glycine | NA-Gly | 6.98 | 362.3 > 76.1 | 25 | 79 | 35 | d8-NA-Gly |
| Dihomo-ϒ-Linolenoyl Ethanolamide | Dihomo GLA EA | 7.02 | 350.3 > 62.1 | 20 | 65 | 36 | d8-AEA |
| 1-Arachidonoyl Glycerol | 1-AG | 7.1 | 379.3 > 287 | 20 | 53 | 19 | d5-2-AG |
| 2-Linoleoyl Glycerol | 2-LG | 7.11 | 355.3 > 263 | 20 | 52 | 18 | d5-2-AG |
| d4-Palmitoyl Ethanolamide | d4-PEA | 7.31 | 304.2 > 62.1 | 20 | 80 | 35 | CUDA |
| 1-Linoleoyl Glycerol | 1-LG | 7.32 | 355.3 > 263 | 20 | 52 | 18 | d5-2-AG |
| Palmitoyl Ethanolamide | PEA | 7.33 | 300.2 > 62.1 | 20 | 80 | 31 | d8-AEA |
| Docosatetraenoyl Ethanolamide | DEA | 7.45 | 376.3 > 62.1 | 20 | 66 | 36 | d8-AEA |
| Oleoyl Ethanolamide | OEA | 7.6 | 326.2 > 62.1 | 20 | 80 | 32 | d8-AEA |

* - See Table I for UPLC conditions. Collision-induced dissociation was performed with 2.3 mTorr nitrogen. Dashed lines indicate separation between mass spectral multiple reaction monitoring functions.

† - Internal Standards (ISTD) - Analytes were corrected for recoveries of listed surrogates. 1-Cyclohexylureido,3-dodecanoic acid (CUDA; and 1-Phenyl 3-Hexadecanoic Acid Urea (PHAU) was introduced immediately prior to analysis and used to quantify surrogate recoveries.

Table 2b. UPLC/Electrospray ionization QTRAP analyte and instrument specific parameters (continued)*

| Analyte | Common Abbreviation | tR (min) | Transition (Da) | Dwell (msec) | Declustering (V) | Collision (V) | ISTD† |
| --- | --- | --- | --- | --- | --- | --- | --- |
| N-Oleoyl Glycine | NO-Gly | 8.14 | 340.2 > 76.2 | 20 | 80 | 26 | d8-NA-Gly |
| 2-Oleoyl Glycerol | 2-OG | 8.2 | 357.3 > 265 | 50 | 52 | 18 | d5-2-AG |
| 1-Oleoyl Glycerol | 1-OG | 8.86 | 357.3 > 265 | 50 | 52 | 18 | d5-2-AG |
| Stearoyl Ethanolamide | SEA | 8.97 | 328.2 > 62.1 | 50 | 80 | 35 | d8-AEA |

* - See Table Ib for UPLC conditions. Collision-induced dissociation was performed with 2.3 mTorr nitrogen. Dashed lines indicate separation between mass spectral multiple reaction monitoring functions.

† - Internal Standards - Analytes were corrected for recoveries of listed surrogates. 1-Cyclohexylureido,3-dodecanoic acid (CUDA; and 1-Phenyl 3-Hexadecanoic Acid Urea (PHAU) was introduced immediately prior to analysis and used to quantify surrogate recoveries.

| Table 2c. Ceramide UPLC/Electrospray ionization QTRAP analyte and instrument specific parameters* |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Analyte | tR (min) | Transition (Da) | Declustering (V) | Collision (V) | ISTD† |
| 1-Cyclohexyl Urea 3-Dodecanoic Acid | 3.0 | 341.3 > 216.2 | 60 | 24 | --- |
| 17:1 Sphingosine | 3.4 | 286.4 > 268.3 | 40 | 15 | CUDA |
| 18:1 Sphingosine | 3.8 | 300.4 > 282.4 | 40 | 21 | 17: 1 Sphingosine |
| 17:1 Sphingosine-1P | 3.7 | 366.4 > 250.3 | 50 | 23 | CUDA |
| 18:1 Sphingosine-1P | 4.0 | 380.4 > 264.4 | 50 | 25 | 17: 1 Sphingosine-1P |
| 18:0 Sphinganine-1P | 4.0 | 382.4 > 266.4 | 50 | 25 | 17: 1 Sphingosine-1P |
| C14 Ceramide | 8.8 | 510.7 > 492.6 | 50 | 21 | C17 Ceramide |
| C16 Ceramide | 9.5 | 538.8 > 264.4 | 55 | 37 | C17 Ceramide |
| C18:1 Ceramide | 9.8 | 564.5 > 546.4 | 60 | 24 | C17 Ceramide |
| C17 Ceramide | 9.9 | 552.8 > 534.5 | 55 | 24 | CUDA |
| C18 Ceramide | 10.3 | 566.7 > 264.4 | 55 | 37 | C17 Ceramide |
| C18 dihydroceramide | 10.6 | 568.7 > 266.4 | 85 | 33 | C17 Ceramide |
| C20 Ceramide | 11.3 | 594.4 > 576.5 | 55 | 21 | C17 Ceramide |
| C24 Ceramide | 13.3 | 650.9 > 264.4 | 55 | 42 | C17 Ceramide |
| C24 dihydroceramide | 13.6 | 652.9 > 266.4 | 55 | 42 | C17 Ceramide |
| C25 Ceramide | 11.3 | 664.9 > 264.4 | 55 | 45 | C17 Ceramide |
| * - Analytes were separated under conditions described in Table Ic. Collision-induced |  |  |  |  |  |
| dissociation was performed with nitrogen at a pressure of 2.3 mTorr. |  |  |  |  |  |
| Dashed lines indicate separation between mass spectral multiple reaction monitoring |  |  |  |  |  |
| functions. |  |  |  |  |  |
| † - Internal Standards (ISTD) - Analytes were corrected for recoveries of listed surrogates. |  |  |  |  |  |
| 1-Cyclohexylureido,3-dodecanoic acid (CUDA) was introduced immediately prior to analysis |  |  |  |  |  |
| and used to quantify surrogate recoveries. |  |  |  |  |  |